Drager Pac 7000 Hydrogen Cyanide Cheat Sheet

Turn On / Turn Off

- Always turn the meter on/off in clean air

- Turn on: Press and hold OK and the display will count down 3, 2, 1.

  - LED, Alarm, and vibrating alarm are activated in sequence

  - The instrument will perform a self test

  - A1 and A2 alarm limits are displayed

- Turn off: Press and hold + and OK button for three seconds

Fresh air calibration

- Press + three times within three seconds

- The meter will beep

- Press + one additional time after that beep

- Press OK TWO times

- Press + to make the code 001

- Press OK two times

- Press OK when the fan icon with the word CAL appears

- The screen will flash and the fresh air calibration is complete when the fan icon is removed

from the screen

- If the fresh air calibration failed, an X icon will appear and the fresh

Air calibration can be attempted again

Calibration: This meter needs to be calibrated every _________ months.

Switch on the backlight

Sensors and Alarms

| Sensor |  | Range | Low Alarm | High Alarm |
| --- | --- | --- | --- | --- |
| HCN | Hydrogen Cyanide | 0 - 50 PPM | 10 | 20 |

During the operation, if allowable measurement range is exceeded of a negative drift occurs the following will appear in the display:

= Too high Concentration

= Negative Drift

Battery

- Type= Lithium-Ion

- When the battery pre-alarm is activated, the audible alarm sounds and the LED blinks and the low battery icon flashes

- You must acknowledge this by pressing OK

- After this first battery pre-alarm, the battery will last for approximately 1 further week and the low battery icon stays lit

- When the main battery alarm is activated, the audible alarm sounds in a repeating pattern of two tones and the LED flashes in the same pattern

- You cannot acknowledge the main battery alarm. The meter will turn off after approximately 10 seconds

TWA, STEL, and Peak Concentration

TWA – Time weighted average (Conc. To not be exceeded in 8 hour work shift for 40 hour work

week)

STEL – Short Term Exposure Limit (Conc. To not be exceeded in 15 minute period)

To show these readings, press OK while monitoring the environment.  The peak concentration icon will appear:          After 10 seconds, the display will return to the monitoring screen

If OK is pressed again before 10 seconds, the TWA icon and concentration will appear.

If OK is pressed again before 10 seconds, the STEL icon and concentration will appear.

CAUTIONS AND NOTES

- Measure pH first; a corrosive will destroy your meter.

SIGNS AND SYMPTOMS BASED UPON EXPOSURE

PPM			SYMPTOMS

- 0.2 to 5			Detectable Odor

- 10				TLV

- 18 – 36			Slight headache symptoms

- 45 – 54			Tolerated for up to 1 hour

- 100				Fatal within 1 hour

- 110 – 135			Fatal within ½ to 1 hour

- 181				Fatal in 10 minutes

- 280				Immediately fatal

NOTES: